## Sheet1
ANNEX II + III: TECHNICAL SPECIFICATIONS + TECHNICAL OFFER / PRILOG II + III: TEHNIČKE SPECIFIKACIJE + TEHNIČKA PONUDA
Contract title: Procurement of laboratory equipment for micro-propagation and laboratory furniture / Naslov ugovora: Nabava laboratorijske opreme za mikropropagaciju i laboratorijskog namještaja Tender reference / Referenca nabave: IPA2007/HR/16IPO/001/040213/RBTC-004/SUP/3.2.4
Column 1-2-3 should be completed by the Grant Beneficiary / Stupac 1-2-3 ispunjava ugovaratelj. Column 4-5 should be completed by the Tenderer / Stupac 4-5 ispunjava ponuditelj. Column 6 is reserved for the Evaluation Committee / Stupac 6 je rezerviran za Odbor za ocjenjivanje.
The Tenderers are requested to complete the template presented below / Ponuditelji trebaju ispuniti obrazac u nastavku dokumenta:
• Column 3 is completed by the Grant Beneficiary and shows the required specifications (not to be modified by the tenderer) / Stupac 3 ispunio je ugovaratelj i definira tražene specifikacije (ponuditelj ne smije mijenjati tražene specifikacije). • Column 4 is to be filled in by the Tenderer and must detail what is offered (for example the words “compliant” or “yes” are not sufficient) / Stupac 4 ispunjava ponuditelj i ponuđena oprema mora biti detaljno opisana (npr. ispunjavanje stupca riječima “u skladu” ili "da" nije dostatno). • Column 5 allows the Tenderer to make comments on its proposed supply and to make eventual references to the documentation / Stupac 5 omogućuje ponuditelju unos komentara vezanih uz ponuđenu opremu i/ili dodavanje reference na dokumente.
The offer must be clear enough to allow the evaluators to make an easy comparison between the requested specifications and the offered specifications / Ponuda mora biti dovoljno jasna da omogući ocjeniteljima jasnu usporedbu između traženih specifikacija i ponuđenih specifikacija.
TECHNICAL SPECIFICATION FOR SUPPLY UNDER LOCAL OPEN PROCEDURE / TEHNIČKE SPECIFIKACIJE ZA NABAVU OPREME U SKLOPU OTVORENOG NATJEČAJNOG POSTUPKA
1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Ref. Nr. / Red. br. | Quantity / Količina | Technical Specification / Tehnički opis stavke | Specifications offered (including brand and/or model) / Ponuđene specifikacije (uključujući brand i/ili model) | Comments / Komentari | Evaluation Committee notes / Bilješke Odbora za ocjenjivanje | Unit Price in HRK* / Jed. cijena u HRK** | Total HRK* / Ukupno HRK**
Item 1. | 3 | LAMINAR FLOW CABINET LAMINARNI KABINET |  |  |  |  | 0
With base stand, working position: sitting, wide min. 1200 mm, max. 1400 mm. Type of protection: protection of samples and work processes. Type of air flow: laminar/horizontal with adjustable air flow velocity. Main filter efficiency: 99,995%, prefilter according to EN 779. Ilumination: fluo-lamp min. 500 lux. UV bactericide lamp min. 30 W integrated in working space. Setting up a connection for a gas burner in working space. Electrical socket on front area of hood. From front protected with a transparent cover with halfway open cover. Protection with safety system for unfavorable conditions. Low level of noise. Built-in materials: powder coated steel. Work surface: polished stainless steel. Delivery with an additional set of pre-filter and main-filter. Acessibility of authorized service. Delivery with installation, validation and user training. User manual in English and Croatian. S postoljem, radni položaj: sjedeći, širina min. 1200 mm, max. 1400 mm. Tip zaštite: zaštita predmeta rada. Tip strujanja zraka: laminarno/ horizontalno uz podesivu brzinu strujanja zraka. Glavni filter efikasnosti: 99,995%, predfilter prema EN 779. Osvjetljenje fluo-lampa min. 500 lux. UV baktericidna lampa min. 30W postavljena na radnu površinu. Na radnom stolu priključak za plinski plamenik. Električna utičnica (230V) na prednjoj strani uređaja. S prednje strane uređaj zaštićen transparentnim poklopcem s mogućnošću preklopa na pola. Zaštita sa sigurnosnim sustavom za nepovoljna stanja. Niska razina buke. Konstrukcija: metalna, plastificirana. Radna površina: polirani nehrđajući čelik. Isporuka dodatnog seta predfiltera and glavnog filtera. Dostupnost ovlaštenog servisa. Isporuka s ugradnjom, ovjerom valjanosti i obukom korisnika. Priručnik za korištenje na engleskom i na hrvatskom jeziku.
Item 2 | 1 | STEAM STERILIZER AUTOKLAV |  |  |  |  | 0
The volume of 110-120 litres, selfstanding, with wheels. Workig temperature 105-123 C°. Automatic work with the possibility of selection of working programme and sterilizing regime. Fast achieving of working pressure and with integrated fast cooling for termally unstable materials. Steam condensator integrated. Continous monitoring of temperature and pressure during working process. Safety functions: safety lock, safety pressure alarm, prevention from overheating, protection of dry heaters and other unfavourable conditions. Built-in materials: polished stainless steel, powder-coated metal housing. Delivery with sterilizing baskets. Acessibility of authorized service. Delivery with installation and user trainning. User manual in English and Croatian. Volumen 110-120 litara, samostojeći, s kotačima. Radna temperatura 105-123 C°. Automatski rad uz mogućnost odabira programa rada i režima sterilizacije. Brzo postizanje radnog tlaka i ugrađeno brzo hlađenje za termolabilne materijale. Kondenzator pare u uređaju. Kontinuirano praćenje temperature i tlaka za vrijeme rada. Sigurnosne funkcije: sigurnosna brava, sigurnosni alarm za tlak, sprečavanje pregrijavanja, zaštita grijača na suho i ostala nepovoljna stanja. Materijal: polirani nehrđajući čelik, metalno kućište izvana plastificirano. Dostaviti s pripadajućim košarama za sterilizaciju. Dostupnost ovlaštenog servisa. Isporuka s ugradnjom i obukom korisnika. Priručnik za korištenje na engleskom i na hrvatskom jeziku.
Item 3 | 1 | DIGITAL pH METER DIGITALNI pH METAR |  |  |  |  | 0
Measurement range (pH) from 0 to 14; measurement parameters "pH", "mV", and "temperature". Measurement accuracy from 0,002 pH to 0,005 pH, electrode with temperature sensor, electrode holder and magnetic stirring involved. Automatic calibration and buffer recognition. Automatic temperature compensation. Auto reed functions. Accessories: electrodes with temperature sensor (2 pcs), electrode holder, magnetic mixer, reference buffer solution (three buffers). Acessibility of authorized service. Delivery with installation and user trainning. User manual in English and Croatian. Mjerno područje( pH) od 0 do 14; parametri mjerenja "pH", "mV" i "temperatura". Točnost mjerenja od 0,002 pH do 0,005 pH. Elektroda s temperaturnim senzorom, držač elektrode i magnetska mješalica uključeni. Automatska kalibracija i prepoznavanje pufera. Automatska kompenzacija temperature. "Auto reed" funkcija. Dodatna oprema: elektrode s temperaturnim senzorom (2 kom), držač elektrode, magnetna mješalica, otopina referentnih pufera (tri pufera). Dostupnost ovlaštenog servisa. Isporuka s ugradnjom i obukom korisnika. Priručnik za korištenje na engleskom i na hrvatskom jeziku.
Item 4 | 1 | DESK pH METER STOLNI pH METAR |  |  |  |  | 0
pH: 0 - 14. Accuracy 0,005 pH; pH GEL probe with temperature sensor, standard. Automatic electrode recognition and calibration data storage. Robust. Delivery in complete with electrode holder, buffer solutions and all accessories for start measuring. with installation and user trainning. User manual in English and Croatian. pH: 0 do 14. Točnost 0,005 pH; pH GEL sonda s temperaturnim senzorom, standardna. Automatsko prepoznavanje elektrode i pohranjivanje podataka kalibracije. Robustan. Isporučiti u kompletu s držačem elektrode, otopinama pufera i svim dodacima potrebnim za mjerenje. Isporuka s obukom korisnika. Priručnik za korištenje na engleskom i na hrvatskom jeziku.
Item 5 | 1 | REFRIGERATOR HLADNJAK |  |  |  |  | 0
Free-standing refrigerator with icebox. Refrigerator volume: 80-125 litres. Icebox volume: 20-45 litres. Energy Class: min. A+. Acessibility of authorized service. Delivery with installation and user trainning. User manual in English and Croatian. Samostojeći hladnjak s ledenicom. Volumen hladnjaka: 80-125 litara. Volumen ledenice: 20-45 litara. Energetska klasa: min. A+. Dostupnost ovlaštenog servisa. Isporuka s ugradnjom i obukom korisnika. Priručnik za korištenje na engleskom i na hrvatskom jeziku.
Item 6 | 1 | DRY STERILIZER SUŠIONIK STERILIZATOR |  |  |  |  | 0
Capacity 100-120 litres. Temperature regime: 40-250 C°. With digital electronic temperature and time controller. With a built-in fan and a safety thermostat. Acessibility of authorized service. Delivery with installation and user trainning. User manual in English and Croatian. Kapacitet 100-120 litara. Temperaturni režim: 40-250 C°. S digitalnim elektronskim programatorom temperature i vremena. S ugrađenim ventilatorom i sigurnosnim termostatom. Potrebna je dostupnost ovlaštenog servisa. Isporuka s ugradnjom i obukom korisnika. Priručnik za korištenje na engleskom i na hrvatskom jeziku.
Item 7 | 1 | DISHWASHER PERILICA POSUĐA |  |  |  |  | 0
Detached, standard, XXL interior, auto programme, self-cleaning filter, energy class min. A+. Acessibility of authorized service. Delivery with installation and user trainning. User manual in English and Croatian. Samostojeća, standardna, XXL unutrašnjost, auto program, samočisteći filter, energetska klasa min. A+. Potrebna je dostupnost ovlaštenog servisa. Isporuka s ugradnjom i obukom korisnika. Priručnik za korištenje na engleskom i na hrvatskom jeziku.
Item 8 | 1 | ANALYTICAL BALANCES ANALITIČKA VAGA |  |  |  |  | 0
Capacity min. 200, max. 250 g. The accuracy of readings 0,1 mg. The reproducibility of 0,2 mg. Stabilization time max. 4 seconds. Automatic calibration, internal, every two hours or when temperature changes. Display in Croatian or English. Platform stainless steel; diameter of 80-100 mm. Overload protection. Acessibility of authorized service. Delivery with installation and user trainning. Delivery with Measurement and Scale Certificate verified by State Office for Metrology. User manual in English and Croatian. Kapacitet min. 200 g, max. 250. Točnost očitanja 0,1 mg. Reproduktivnost 0,2 mg. Vrijeme stabilizacije max. 4 sekunde. Automatska kalibracija, interna, svakih dva sata ili kod promjene temperature. Display na hrvatskom ili engleskom jeziku. Platforma inox; promjera 80-100 mm. Zaštita od preopterećenja. Potrebna je dostupnost ovlaštenog servisa. Isporuka s ugradnjom i obukom korisnika. Isporučiti s Certifikatom o umjeravanju i ovjerom vage od strane Državnog zavoda za mjeriteljstvo. Priručnik za korištenje na engleskom i na hrvatskom jeziku.
Item 9 | 1 | OVERHEAD STIRER ROTACIONA MJEŠALICA |  |  |  |  | 0
Purpose: mixing viscous media in sharp enviroment, volume of mixing media 10-20 liltres. Design of direct axe should provide regulation of agitator height. Hermetically sealed set. Minimum of two mixing speeds. Delivery with stand, klemm and propeller bar. Acessibility of authorized service. Delivery with installation and user trainning. User manual in English and Croatian. Namjena: miješanje viskoznog medija u oštrom okruženju, volumen medija za miješanje 10-20 litara. Dizajn direktne osovine treba omogućiti regulaciju visine mješalice. Hermetički zatvoreno kućište. Minimalno dvije brzine miješanja. Isporučiti komplet sa stalkom, klemom i šipkom sa propelerom. Potrebna je dostupnost ovlaštenog servisa. Isporuka s ugradnjom i obukom korisnika. Priručnik za korištenje na engleskom i na hrvatskom jeziku.
Item 10 | 1 | DESK GAS STOVE STOLNO PLINSKO KUHALO |  |  |  |  | 0
Detached, with two burners: - regular burner (min. 1.0 kW, max. 1,75 kW) - strong burner (min. 1.75 kW, max. 3,0 kW) With fuse ignition of gas burners. Acessibility of authorized service. Delivery with installation and user trainning. User manual in English and Croatian. Samostojeće, s dva plamenika: - normalni plamenik (min. 1,0 kW, max 1,75 kW) - jaki plamenik (min. 1,75 kW, max. 3,0 kW) S osiguračima paljenja plinskih plamenika. Potrebna je dostupnost ovlaštenog servisa. Isporuka s ugradnjom i obukom korisnika. Priručnik za korištenje na engleskom i na hrvatskom.
Item 11 | 3 | DIGITAL THERMOMETER AND HYDROMETER DIGITALNI TERMOMETAR I HIGROMETAR |  |  |  |  | 0
Ambient with internal and external sensors, and possibility of recording the lowest and the highest temperatures measured. Temperature range from (-10) to (+60) C°. Humidity range from 10% to 99%. Resolution 0.1 C°, 1% RH. Acessibility of authorized service. Delivery with user trainning. User manual in English and Croatian. Ambijentalni s internim i eksternim senzorima, i mogućnošću bilježenja najniže i najviše izmjerene temperature. Temperaturni raspon od (-10) do )+60) C°. Vlaga od 10% do 99%. Razlučivost 0,1 C°, 1% relativne vlage. Potrebna je dostupnost ovlaštenog servisa. Isporuka s obukom korisnika. Priručnik za korištenje na engleskom i na hrvatskom jeziku.
Item 12 | 1 | DESK EC METER STOLNI EC METAR |  |  |  |  | 0
The device for measuring the electrical conductivity. Measurement range 0-20.0 mS. Accuracy 0.5%, Electrode: standard epox. Flexible working and storage of calibration data, robust. Acessibility of authorized service. Delivery with installation and user trainning. User manual in English and Croatian. Uređaj za mjerenje elektroprovodljivosti. Raspon mjerenja 0-20,0 mS. Točnost 0,5%. Elektroda: stadardna epox. Fleksibilan rad i pohrana podataka kalibracije, robustan. Potrebna je dostupnost ovlaštenog servisa. Isporuka s ugradnjom i obukom korisnika. Priručnik za korištenje na engleskom i na hrvatskom jeziku.
Item 13 | 1 | DESK ILLUMINATED MAGNIFIER STOLNO POVEĆALO S RASVJETOM |  |  |  |  | 0
Magnifieing min. 3x, with fluorlescent light (min. 9 W), tightend on mobile holder. Povećanje min. 3x, s fluorescentnim osvjetljenjem (min. 9 W), učvršćeno na pokretnom držaču.
Item 14 | 2 | LABORATORY SINK LABORATORIJSKI SUDOPER |  |  |  |  | 0
Dimensions approx. 1500x600x900mm (WxDxH). Work surface: polypropylene with two polypropylene cannals approx. 500x400mm. Laboratory hot/cold water mixer, table setup. Cabinet with solid doors made of humidity resistant material RAL 7047 and the edges are protected by ABS ribbon. With drain. In accordance with HRN D.E2.104.HRN D.E2.105. Acessibility of authorized service. Delivery with installation and user trainning. Dimenzije cca. 1500x600x900mm (ŠxDxV). Radna površina polipropilen s dva polipropilenska korita cca 500x400mm. Laboratorijska mješalica topla/hladna voda, stolna montaža. Ormarić s punim vratima, izrađen od vlagootpornog materijala, RAL 7047, a rubovi zaštićeni ABS trakom. S priključkom za odvod u kanalizaciju.Sukladno HRN D.E2.104.HRN D.E2.105.Potrebna je dostupnost ovlaštenog servisa. Isporuka s ugradnjom i obukom korisnika.
Item 15 | 5 | WALL MOUNTED DRYER FOR LABORATORY DISHES ZIDNI STALAK ZA SUŠENJE POSUĐA |  |  |  |  | 0
Dimensions approx. (50-70) cm x (50-70) cm. Material: humidity resistant. Delivery with installation (i.e. to be wall mounted) and user trainning. Dimenzije cca (50-70) cm x (50-70) cm. Materijal otporan na vlagu. Isporuka s ugradnjom (tj. postavljanjem na zid) i obukom korisnika.
Item 16 | 1 | FUME HOODS DIGESTOR |  |  |  |  | 0
Dimensions approx. 1200x950x2550 mm. Work surface made of monolite tehnical ceramics with mini-sink, cold water, draining, gass, side doors with safety glass, 2x230 V socket, ventilation connection, propeler motor. In accordance with HRN EN 14175-2. Acessibility of authorized service. Delivery with installation and user trainning. User manual in English and Croatian. Dimenzije cca 1200x950x2550 mm. Radna površina od monolitne tehničke keramike s mini koritom, hladna voda, odvod, plin, okomito-vodoravna pomična vrata sa sigurnosnim staklom, 2x230 V utičnica, priprema za spoj na ventilaciju, motor s ventilatorom.Sukladno HRN EN 14175-2. Potrebna je dostupnost ovlaštenog servisa. Isporuka s ugradnjom i obukom korisnika. Priručnik za korištenje na engleskom i na hrvatskom jeziku.
Item 17 | 6 | LABORATORY TABLE, TYPE A LABORATORIJSKI STOL, TIP A |  |  |  |  | 0
Dimensions 1500x750x750mm (WxDxH), minor dimension deviations (up to ±2%) acceptable. Doubled metal construction protected with epoxy coating, without under counter elements, work surface: compact. Carrying capaciity 180-220 kg/m. Acessibility of authorized service. Delivery with installation and user trainning. Dimenzije 1500x750x750 (ŠxDxV), manja odstupanja dimenzija (up to ±2%) prihvatljiva. Metalna, ojačana konstrukcija zaštićena epoxy premazom, bez podpultnih elemenata, radna površina: compact.Nosivost 180-220 kg/m. Potrebna je dostupnost ovlaštenog servisa. Isporuka s ugradnjom i obukom korisnika.
Item 18 | 2 | LABORATORY TABLE, TYPE B LABORATORIJSKI STOL, TIP B |  |  |  |  | 0
Dimensions approx. 1800x750x900 mm (WxDxH), working position: standing. Metal construction with epoxy coating, under counter combo of lockers and cabinets, made of humidity resistant chipboard. All edges protected with ABS ribbon. Drawers with metal bearings. Settin up locking doors, great capacity and stability. Quality of work surface material: compact. Gass, water and electricity connection neccessery on the work surface; perform drainage trough. Acessibility of authorized service. Delivery with installation and user trainning. Dimenzije cca. 1800x750x900 mm (ŠxDxV), pozicija za rad: stojeći. Metalna konstrukcija s epoxy premazom. Podpultni elementi kombinacija ormarića i ladičara, izrađeni od vlagootpornog iverala.Svi rubovi zaštićeni ABS trakom. Ladice metalne s ležajevima. Postaviti bravice za zaključavanje, velika nosivost i stabilnost. Kvaliteta materijala radne površine: compact. Na radnu površinu postaviti priključak za plin, vodu i struju i izvesti malo izljevno korito. Potrebna je dostupnost ovlaštenog servisa. Isporuka s ugradnjom i obukom korisnika.
Item 19 | 1 | LABORATORY TABLE, TYPE C LABORATORIJSKI STOL, TIP C |  |  |  |  | 0
Dimensions approx. 2700x750x900 mm (WxDxH), working position: standing. Metal construction with epoxy coating. Under counter combo of lockers and cabinets, made of humidity resistant chipboard. All edges are protected with ABS ribbon. Setting up locking doors, great capacity and stability. Great capacity and stability. Quality of work surface material: compact. Gass, water and electricity connection neccessery on the work surface; perform drainage trough. Acessibility of authorized service. Delivery with installation and user trainning. Dimenzije cca 2700x750x900 mm (ŠxDxV), za stojeći rad. Metalna konstrukcija s epoxy premazom. Podpultni elementi kombinacija ormarića i ladičara, izrađeni od vlagootpornog iverala. Svi rubovi zaštićeni ABS trakom. Ladice metalne s ležajevima. Postaviti bravice za zaključavanje. Velika nosivost i stabilnost. Kvaliteta materijala radne površine: compact. Na radnu površinu postaviti priključak za plin, vodu i struju i izvesti malo izljevno korito. Potrebna je dostupnost ovlaštenog servisa. Isporuka s ugradnjom i obukom korisnika.
Item 20 | 3 | PENDENT LABORATOY CABINET VISEĆI LABORATORIJSKI ORMARIĆ |  |  |  |  | 0
Dimensions approx. 840x330x380 mm (WxDxH), made of waterproof chipboard (resistant to humidity), with solid protected doors. All edges protected by a ABS ribbon. Acessibility of authorized service. Delivery with installation and user trainning. Dimenzije cca 840x330x380 mm (ŠxDxV), izrađen od vodootpornog iverala, s punim otklopnim vratima. Svi rubovi zaštićeni ABS trakom. Potrebna je dostupnost ovlaštenog servisa. Isporuka s ugradnjom i obukom korisnika.
Item 21 | 1 | SAFETY STORAGE CABINET FOR CHEMICALS SIGURNOSNI ORMAR ZA KEMIKALIJE |  |  |  |  | 0
Vented safety storage cabinet for acids and alkalies. Dimensions approx. 600x500x720 mm. Located under fume hood; delivered prepared for connection. Acessibility of authorized service. Delivery with installation and user trainning. Ventilirani sigurnosni ormar za kiseline i lužine. Dimenzije cca 600x500x720 mm. Smješten ispod digestora; isporučiti s pripremom za spajanje. Potrebna je dostupnost ovlaštenog servisa. Isporuka s ugradnjom i obukom korisnika.
Item 22 | 9 | LABORATORY STOOL LABORATORIJSKA STOLICA |  |  |  |  | 0
Ergonomical stool, with backrest, made of polyurethane, on wheels with hydraulics for adjusting the height. Material adjusted for maintenance and dezinfection. Acessibility of authorized service. Delivery with assembly and user training. Ergonomska, s naslonom, od poliuretana, s kotačima i hidraulikom za podešavanje visine. Materijal podesan za održavanje i dezinfekciju. Potrebna je dostupnost ovlaštenog servisa. Isporuka s montažom i obukom korisnika.
Item 23 | 5 | TRANSPORT TROLLEY TRANSPORTNA KOLICA |  |  |  |  | 0
For safe transport of jars in laboratory on wheels (the possibility of fixing two wheels). Dimensions approx. 400x650x820mm, with two shelves enlarged capacity, min. 20 kg per shelf. Material adjusted for maintenance and dezinfection. Acessibility of authorized service. Služe za siguran transport staklenki u laboratorijskom prostoru, na kotačima (mogućnost fiksiranja dva kotača). Dimenzije cca 400x650x820 mm, s dvije police povećane nosivosti, min. 20 kg po polici. Materijal izvedbe podesan za održavanje i dezinfekciju. Potrebna je dostupnost ovlaštenog servisa.
Item 24 | 1 | WEIGHING TABLE WITH SOLID ARTIFICIAL STONE STOL ZA VAGANJE S KAMENOM PLOČOM |  |  |  |  | 0
Dimensions approx. 900x600x750 mm. Acessibility of authorized service. Dimenzije cca 900x600x750 mm. Potrebna je dostupnost ovlaštenog servisa.
TRAINING OBUKA OSOBLJA 1 day training on how to use the equipment is to be provided for the 3 staff of High-school M.A. Reljković Slavonski Brod in the place of delivery. The cost of training is included in the unit price. Potrebno je isporučiti 1-dnevnu obuku za rad s isporučenom opremom za 3 djelatnika SŠ M.A. Reljkovića Slavonski Brod, na mjestu isporuke. Trošak obuke uključen je u jediničnu cijenu.
COMMERCIAL WARRANTY KOMERCIJALNA GARANCIJA required for items / se zahtijeva za stavke Item 1 - Laminar flow cabinet / Laminarni kabinet Item 2 - Steam sterilizer / Autoklav Item 3 - Digital pH meter / Digitalni pH metar Item 4 - Desk pH meter / Stolni pH metar Item 6 - Dry sterilizer / Sušionik sterilizator Item 8 - Analytical balances / Analitička vaga Item 9 - Overhead stirer / Rotaciona mješalica Item 12 - Desk EC meter / Stolni EC metar for the period of 2 years counting from the date of the expiration of the standard warranty. / za razdoblje od dvije godine od dana isteka standardne garancije.
RESPONSE TIME VRIJEME ODAZIVA On-site response time within 24 hours during 2 years after provisional acceptance. / Vrijeme odaziva na lokaciju do 24 sata za razdoblje od 2 godine od dana privremenog prihvata.
Total HRK* / Ukupno HRK* | [6 empty] | 0
* Remark: Do not fill in shaded area / Napomena: Ne ispunjavati osjenčani dio.
** Remark: Prices are quoted without VAT / Napomena: Cijene su bez PDV-a.
I hereby declare that, should we win this tender, we shall supply, deliver, install, commision and test these goods FCA to High School M. A. Reljkovic from Slavonski Brod, Ivana Cankara 76, 35000 Slavonski Brod and train the staff no later than 20th of November 2013 / U slučaju prihvaćanja naše ponude izjavljujemo da ćemo navedenu robu nabaviti, isporučiti, ugraditi, pustiti u pogon i testirati FCO u Srednjoj školi M. A. Reljkovića Slavonski Brod, Ivana Cankara 76, 35000 Slavonski Brod i obučiti osoblje najkasnije do 20. studenoga 2013.
Name of the Firm / Naziv tvrtke:
Name of the Firm's authorised representative / Ime ovlaštenog predstavnika tvrtke:
Signature of the Firm's authorised representative / Potpis ovlaštenog predstavnika tvrtke:
Date and Firm's stamp / Datum i pečat tvrtke: